Abstract Reasoning Experiment - Procedure

General Preparation

- Instructions to participants:

  - Before the day of the expriment:

    - No hair product

    - No eye makeup

    - If corrected-vision, wear contact lenses instead of glasses if possible

  - Try to refrain from moving, clenching their jaw, and blinking as much as possible

  - During breaks, take as long as needed but try to keep head still and not move too much

- Measure screen to chin-rest distance

- Check tape marks on location of screen and EyeLink camera

- Check screen resolution and refresh rate

- File naming:

  - Pattern: “cp<participant ID><session number>”

  - using two digits (leading zero) for both Participant ID  and Session number

  - e.g., participant 1, session 1: “cp0101”

First Steps

- If it is the participant’s first session: demographics form

- Informed consent form

- Turn on all three experimenter’s PCs and the participant’s PC

- Turn on light, decrease brightness to minimum. Make sure the neons are not flickering, otherwise try to increase the brightness a bit until they stop.

- Plug in the Eye tracker

- connect the battery to the EEG amplifier and turn it on

- turn off the speakers

- EEG computer:

  - Launch ActiView

  - Check the battery level, swap the battery if needed

  - Select configuration file:

    - Click on "About / configure" tab

    - Click on “Load CFG File”

  - Select "Michael2.cfg"

  - Press start on upper left corner

- EyeLink computer:

  - click on "Tracker"

  - Load the correct configuration

- Ask participant to put phone / electronic devices away

Setting up EEG – Part 1

- Measure the head’s circumference and select the right cap size

| Small / Medium | 52-56 cm | Yellow & Red |
| --- | --- | --- |
| Medium | 54-58 cm | Red |
| Medium / Large | 56-60 cm | Blue & Red |
| Large | 58-62 cm | Blue |

- Fit the cap on the head, get ears out if more comfortable for participant

- Make sure the cap is aligned:

  - Measure the distance between the inion and nasion, Cz should be at the midpoint

  - Measure the distance between the tragi (protrusion on front of the ear), Cz should be at the midpoint

- Fasten the cap’s strap and ensure it is relatively tight

- Apply conductive gel to EEG cap

  - Fill a clean syringe with conductive gel

  - Gently part the hair beneath each electrode’s hole by slowly rotating the tip of  the syringe on the scalp

  - Slowly inject a small amount of gel while gradually withdrawing the syringe

- Place the eye electrodes

- Check the electrodes activity on the monitor and readjust the cap / gel if necessary

  - Note any problematic electrodes or unusual circumstances in a “notes.txt” file in the session’s folder

Setting up Eye-tracking – Part 1

- Launch the experiment’s script:

  - open the command line and cd to the experiment’s folder: C:\Users\topuser\Documents\ChrisPinier\abstract_reasoning\experiment-Lab

  - run “poetry run python experiment.py”

- Perform dominant eye test

- Fill out the dialogue box:

- Set up chin-rest

- Measure eye to screen distance, record it in the dialogue box (in mm)

- Adjust camera focus

- Click on the pupil of the dominant eye using the mouse

- Press A multiple times to use the auto-threshold until you get stable CR value

Setting up the EEG – Part 2

- Click on "Start File".

- Fill out the fields (copy filename), ignore Error 7003 -> click on “continue”

- Navigate to the data folder: C:\Data\ChrisPinier\abstract-reasoning

- Enter the filename using the pattern described in the “General Preparation” section

- Click on “Paused” to start saving data, the button above will turn green and show “Saving”

Launching the Experiment

- Press C for calibration, move mouse cursor away, press enter to start

- Press V for validation, move mouse cursor away, press enter to start, press enter to continue once calibration done

- Press O to start the experiment, move mouse cursor away

End of experiment

- On the “EEG computer:

  - In ActiView:

    - click on “Pause File” (upper right corner)

    - click on “Stop” (upper left corner)

  - Copy the EEG (.bdf) file to the USB stick

- On the “main computer”:

  - Copy the behavioral data files (practice and experiment; .csv) and the eye tracking files (.edf & .asc) to the USB stick

- Remove the electrodes, put them in the plastic bucket, making sure they are not touching anything metallic

- Remove the cap

- Provide a towel and shampoo to the participant, guide them to the bathroom

- Clean the electrodes, hang them on the wall, electrodes inward, connector outward

- Clean the cap, hang it with the electrodes

Lab Information
Technical specifics

All PC’s standard installed software:

- Windows 10 version 19.03 64bit

- Psychopy

- Eye Link software in Eye Link labs

Hardware:

- Eye Link 1000 (in Eye Link lab)

- 144 Hz stimulus monitor (LG 27GL850 27″ resolution 2560 x 1440)

- Biosemi Actiview, 64 channel Biosemi Active II EEG amplifier (if available)

- High performance stimulus PC to allow real-time data analysis